# Wergonic Device Manager

User Manual

Version 2.0 | February 2026

# 1. Introduction

The Wergonic Device Manager is a web application for configuring and monitoring Wergonic haptic feedback devices. These devices monitor arm or trunk elevation angles and provide vibration alerts when posture thresholds are exceeded.

Key Features:

- Connect to one or two devices simultaneously (ARM and TRUNK)

- Configure vibration intensity, thresholds, and feedback settings

- Calibrate devices with a single button press

- Monitor live angle data in real-time

- Download session data files (firmware v4.0+)

- Set date/time and session subject names (firmware v4.0+)

# 2. Requirements

## 2.1 Browser Requirements

The webapp uses Web Bluetooth API, which requires:

- Google Chrome (recommended) or Microsoft Edge

- Bluetooth enabled on your computer or mobile device

- HTTPS connection (required for Web Bluetooth)

Note: Safari and Firefox do not support Web Bluetooth.

## 2.2 Firmware Versions

The webapp supports both legacy and modern firmware versions:

| Version | Features | Limitations |
| --- | --- | --- |
| Legacy (<4.0) | Calibration, thresholds, vibration intensity, feedback on/off, filter on/off | No SD card logging, no datetime, no session names |
| v4.0+ | All legacy features + SD card logging, datetime setting, session names, file download | None |

# 3. Getting Started

## 3.1 Accessing the Application

- Open Google Chrome or Microsoft Edge

- Navigate to the Wergonic Device Manager URL

- If prompted, enter the access password

- Ensure Bluetooth is enabled on your device

## 3.2 Interface Overview

The application is divided into several panels:

| Panel | Description |
| --- | --- |
| Device Connections | Connect/disconnect ARM and TRUNK devices, calibrate |
| Device Configuration | Set thresholds, intensity, feedback, filter, datetime, session name |
| Live Data | Real-time angle and zone display for connected devices |
| Session Files | List and download session data files (v4.0+ only) |
| Console | Debug mode - send raw commands and view responses |

# 4. Typical User Journey

Follow these steps for a typical session with Wergonic devices:

## Step 1: Connect Device(s)

- Turn on your Wergonic device(s)

- In the Device Connections panel, click "Connect ARM" or "Connect TRUNK"

- A browser popup will show available devices (named "Wergonic-XX")

- Select your device and click "Pair"

- The status dot will turn green when connected

Tip: You can connect both ARM and TRUNK devices simultaneously for dual monitoring.

## Step 2: Configure Settings (Optional)

Before calibration, you may adjust the following settings:

Vibration Intensity: Choose Weak, Medium, or Strong. This can be changed during a session without recalibration.

Thresholds: Adjust the Yellow (warning) and Red (alert) angle thresholds using the sliders. ARM and TRUNK have separate threshold controls with different defaults.

| Device | Yellow Default | Red Default |
| --- | --- | --- |
| ARM | 30° | 60° |
| TRUNK | 20° | 45° |

Complementary Filter: Enable for smoother angle readings (recommended). Disable for raw accelerometer-only readings.

Feedback: Enable to activate vibration feedback when thresholds are exceeded.

## Step 3: Set DateTime and Session Name (v4.0+ only)

For firmware v4.0+, you can set the device's clock and session subject name for better data logging:

- Click "Set Current Time" to sync the device with your computer's clock

- Enter a subject name (max 16 characters) and click "Set Name"

## Step 4: Calibrate

Important: Calibration sets the device's reference "zero" position. The user should be in their neutral posture when calibrating.

- Have the user assume their neutral/correct posture

- Click "Calibrate Connected Devices" to calibrate all connected devices at once

- Or use the "Calibrate" button in Device Configuration to calibrate the selected target device

- You will see "ACK:C:Calibrated" confirmation (v4.0+ firmware)

## Step 5: Monitor Live Data

After calibration, the Live Data panel shows real-time readings:

- Angle: Current elevation angle in degrees

- Zone: GREEN (safe), YELLOW (warning), or RED (alert)

- The device will vibrate according to the zone if feedback is enabled

## Step 6: Standalone Operation

Important: Once calibrated and feedback is started, the device can work standalone without the webapp!

- After calibration, you can disconnect the device from the webapp

- The device will continue monitoring angles and providing vibration feedback

- Session data is logged to the SD card (v4.0+ firmware)

- Reconnect later to download session files or change settings

## Step 7: End Session

- Reconnect to the device if disconnected

- Click "Stop & Save Session" to save data to SD card

- Use the Session Files panel to list and download session data

# 5. Command Reference

The following commands are available. Most are sent automatically via the UI, but can also be sent manually in Debug Mode.

| Command | Action | Response (v4.0+) | Firmware |
| --- | --- | --- | --- |
| C | Calibrate device | ACK:C:Calibrated | All |
| A | Set device type to ARM | ACK:A:ARM | All |
| B | Set device type to TRUNK | ACK:B:TRUNK | All |
| S / M / W | Set intensity Strong/Medium/Weak | ACK:S:Strong | All |
| F / N | Enable/Disable feedback | ACK:F:ON / ACK:N:OFF | All |
| X / Z | Enable/Disable filter | ACK:X:ON / ACK:Z:OFF | All |
| K | Stop & Save session | ACK:K:Session saved | v4.0+ |
| T:YYYYMM... | Set datetime | ACK:T:2026-02-02 21:30 | v4.0+ |
| N:name | Set session subject name | ACK:N:SubjectName | v4.0+ |
| ? | Get device status | Status printout | All |
| D | List SD card files | FILES:BEGIN... | v4.0+ |
| R:filename | Download file | STREAM:BEGIN... | v4.0+ |

# 6. Troubleshooting

| Problem | Solution |
| --- | --- |
| Device not appearing in scan | Ensure device is on and in range. Check Bluetooth is enabled. Try refreshing the page. |
| "GATT operation already in progress" | Wait a moment and retry. The webapp queues commands to prevent this. |
| Device shows as "Arduino" | Update firmware to latest version which uses "Wergonic-XX" naming. |
| No live data after calibration | Angle data requires calibration. Ensure device type is set and calibration completed. |
| Features grayed out | Some features require firmware v4.0+. Check firmware version in device info. |
| Thresholds not applying | Click "Apply to ARM" or "Apply to TRUNK" after adjusting sliders. |

# 7. Debug Mode

Enable Debug Mode using the toggle in the header to access advanced features:

- Console Panel: Send raw commands and view all BLE communication

- Show All Devices: BLE scan shows all nearby devices, not just Wergonic

- Command Responses: View ACK/ERROR responses from the device

# 8. Data Storage (v4.0+)

Firmware v4.0+ stores session data on an SD card. Each session creates two files:

## 8.1 Session Data File (CSV)

Filename: session.csv, s1.csv, s2.csv, etc.

Format: Comma-separated values with the following columns:

| Column | Description |
| --- | --- |
| elapsed_ms | Milliseconds since session start (calibration) |
| timestamp | Human-readable timestamp (HH:MM:SS.mmm) |
| angle | Current angle in degrees (float) |
| feedback | Feedback triggered: 0 = no, 1 = yes |
| zone | Current zone: green, yellow, or red |

Example CSV content:

elapsed_ms,timestamp,angle,feedback,zone

0,00:00:00.000,15.2,0,green

100,00:00:00.100,32.5,1,yellow

200,00:00:00.200,45.8,1,yellow

300,00:00:00.300,62.1,1,red

## 8.2 Metadata File

Filename: session_meta.txt, s1_meta.txt, etc.

Contains session configuration and settings:

| Field | Description |
| --- | --- |
| subject | Subject name set via N: command |
| device_type | ARM or BACK |
| device_id | Device serial number |
| start_date | Session start date (YYYY-MM-DD) |
| start_time | Session start time (HH:MM:SS) |
| threshold_yellow | Yellow threshold in degrees |
| threshold_red | Red threshold in degrees |
| filter_enabled | true/false - complementary filter status |
| gyro_favoring | Filter coefficient (default 0.96) |

Example metadata content:

[Session Metadata]

subject=John_Doe

device_type=ARM

device_id=42

start_date=2026-02-02

start_time=14:30:00

threshold_yellow=30.0

threshold_red=60.0

filter_enabled=true

gyro_favoring=0.96

## 8.3 Downloading Session Files

- Connect to the device in the webapp

- Go to the Session Files panel

- Select ARM or TRUNK device

- Click "List Files" to see available sessions

- Click "Download" to save a file to your computer

- Click "Meta" to view the metadata for a session

# 9. Appendix: Zone Colors

| Zone | Condition | Vibration |
| --- | --- | --- |
| GREEN | Angle < Yellow threshold | None |
| YELLOW | Yellow <= Angle < Red | Double pulse (60% intensity) |
| RED | Angle >= Red threshold | Continuous (80% intensity) |

— End of Manual —